MEMBRANE FILTER INTEGRITY TEST METHODS NYLON MEMBRANE

Non-Sterile 13 mm, 25 mm, 47 mm, 142 mm & 293 mm with 0.1 µm, 0.2 µm 0.45 µm and 0.8 µm Nylon66 (Hydrophilic) Membrane

www.axivalab.com

- INTRODUCTION

  - Introduction

Integrity Test Method is designed to provide a basic understanding of the methods used to qualify Axiva membrane Filter units. Integrity testing sterilizing filters is a fundamental requirement of critical process filtration applications in the pharmaceutical industry. FDA Guidelines require integrity testing of filters used in the processing of sterile solutions such as large volume parenterals (LVPs) and small volume parenterals (SVPs). The FDA also requires corresponding testing documentation be included with batch product records.

  - Product Description

Nylon membrane Filter

Axiva membrane Filter unit exists in pore rating of 0.1μm, 0.2μm, 0.45μm and 0.8μm in membrane diameter of 13 mm, 25 mm, 47 mm, 142 mm and 293 mm Nylon membrane. It removes microorganisms, particles, precipitates, and un-dissolved solid particles / high particulate solution from aqueous or organic solutions. They are non-pyrogenic, non-toxic and single use products. Nylon filters are manufactured using a unique impregnate process that eliminates cracking, tearing, breaking and distortion during use.

Axiva membrane filter have low extractable which assure them for ideal solvent filtration, sample preparation, and mobile phase filtration for HPLC applications. Therefore they are more suitable for use in Analytical Research applications. They are chemically inert and none of the membrane is treated by any means of adhesive, detergents or surfactants makes them ideal for filtering more aggressive solutions. It is designed to meet speed up and increase sample volume while reducing back pressure caused by the blocking of an unprotected membrane. Typical research laboratory applications include:

- Sample filtration for HPLC Analysis and LC-MS Analysis

- Filtration of aqueous as well as organic solvent filtration

- Suitable for method development processes with high product recovery

- Applicable to sterile filtration of solution of interest, buffer solutions and salt solutions

- Clarifications and sterile filtration of biological solutions, tissue culture media, additives, buffers, and water

  - Integrity Test Summary

The integrity test is the End User’s method to confirm the structural integrity of a membrane filter before and after use. There are two classifications of integrity testing: destructive and nondestructive. Axiva’s practice is to perform destructive testing as a lot release criteria on samples from each manufacturing lot of all fabricated sterilizing-grade filter products, and nondestructive testing on each sterilizing-grade filter prior to release to insure its integrity.

- Destructive Integrity Test: Destructive challenge testing is the best way to determine a sterilizing filter’s ability to retain bacteria. Bacterial challenge testing provides assurance that the membrane and fabricated device meet the critical performance criteria of a sterilizing filter. The test is performed on a statistical sample from numerous production lots and fabricated devices produced.

Bacterial Challenge Test

This test confirms the bacterial retentive properties of the Axiva Nylon membrane filter device in accordance with the recognized definition of a sterilizing grade filter membrane. The test is performed according to HIMA Document No.3, Vol. 4 (April 1982) “Microbiological Evaluation of Filters for Sterilizing Liquids”.

- Non-destructive Integrity Test: Non-destructive integrity testing is done in order to correlate the results of the destructive Bacteria Challenge Test. This test provides assurance that the filter will perform as intended. During production of sterile product, the filter should be subjected to such an integrity test before and after filtration. This is done to ensure that the filter meets specification, is properly installed and intact during filtration, and to confirm the rating of the filter.

Following tests are used to determine the integrity of a membrane filter: Bubble Point Test, Burst Pressure Test, and Water Flow Rate Test. All tests are based on the same physics, the flow of an air through a liquid-wetted membrane under applied air pressures.

- Bubble Point Test

The most widely used non-destructive integrity test is the bubble point test. This test consists of measuring the minimum air pressure required to remove the water retained by surface tension in the largest pores of a fully wetted membrane. This pressure (the bubble point) is dependent upon the pore diameter and the surface tension of the wetting fluid. The filter unit is first wet with water, and then the air pressure is slowly increased until a steady stream of bubbles is observed downstream of the filter holder. This pressure is recorded as the bubble point, applied with following formula:

Where,

P = bubble point pressure 		d = pore diameter
k = shape correction factor 		cos θ = liquid-solid contact angle
Σ = surface tension

- Burst Pressure Test

This test consists of measuring the minimum air pressure required to burst of a fully wetted membrane, used as a measure of resistance to rupture. Burst strength depends largely on the tensile strength and extensibility of the material. It is expressed commonly in pounds per square inch (psi). A high burst pressure assures that the system maintains its integrity and does not leak during usage or sample may bypass the membrane.

- Water Flow Rate Test

The purpose of this test is to determine the Axiva membrane Filter meets the minimum specified water flow rate. A sufficient amount of various production lots were taken into account and each filter units were connected to a pressure vessel containing water. Water is passed through the 25 mm S.S. filter holder through a Pressure Vessel at specified 10 psi pressure and collected in a graduate cylinder for 60 seconds.

### AXIVA QUALIFICATION TEST METHODS & RESULTS

  - Destructive Integrity Test

Bacterial Challenge Test

One filter unit from each qualification lot is challenged with a suspension of minimum concentration of 107 organisms per square centimeter of filter surface area of Brevundimonas diminuta for 0.2 µm & Serratia marcescens for 0.45 µm. Each filtrate that passes through Axiva’s membrane Filter is collected separately and placed on agar plate for incubation of 7 days at 30 ± 20C.  There is no growth found in the filtrate.

  - Non-destructive Integrity Test

- Bubble Point Test

This test consists of measuring the minimum air pressure required to remove the distilled water retained by surface tension in the largest pores of a fully wetted membrane. This pressure (the bubble point) is dependent upon the pore diameter and the surface tension of the wetting fluid. The filter unit is first wet with distilled water, and then the air pressure is slowly increased until a steady stream of bubbles is observed downstream of the filter holder. This pressure is recorded as the bubble point.

Axiva Nylon6, 6 membrane filter Bubble Point Lot Releasing criteria

| 0.1 µm | ≥60 psi |
| --- | --- |
| 0.2 µm | ≥52 psi |
| 0.45 µm | ≥30 psi |
| 0.8 µm | ≥16 psi |
| 1.2 µm | ≥15 psi |

Note

The Bubble Point Test results are influenced by the nature of the wetting medium. The bubble point values listed in this validation guide are for Axiva’s membrane filters wetted with water at 20°C. It should be noted, that a variation of the test conditions such as temperature, wetting liquid or type of air pressure may require a different integrity test limit related to those mentioned above.

- Burst Pressure Test

This test consists of measuring the minimum air pressure required to burst of a fully wetted membrane, used as a measure of resistance to rupture. Burst strength depends largely on the tensile strength and extensibility of the material.

Axiva Nylon6, 6 membrane filter Burst Pressure Lot Releasing criteria

| 0.1 µm | ≥13 psi |
| --- | --- |
| 0.2 µm | ≥14 psi |
| 0.45 µm | ≥12 psi |
| 0.8 µm | ≥10 psi |
| 1.2 µm | ≥12 psi |

- Water Flow Rate Test

The purpose of this test is to determine the Axiva membrane Filter meets the minimum specified water flow rate. A sufficient amount of various production lots were taken into account and each filter units were connected to a pressure vessel containing water. Water is passed through the 25 mm S.S. filter holder through a Pressure Vessel at specified 10 psi pressure and collected in a graduate cylinder for 60 seconds.

Axiva Nylon6, 6 membrane filter Water Flow Rate Lot Releasing criteria

| 0.1 µm | ≥ 25 mL/min |
| --- | --- |
| 0.2 µm | ≥ 30 mL/min |
| 0.45 µm | ≥65 mL/min |
| 0.8 µm | ≥130 mL/min |
| 1.2 µm | ≥170 mL/min |